AUTOPROMPTTM CONVERTER

MODEL 2411

USER'S MANUAL

RUSKA INSTRUMENT CORPORATION

10311 WESTPARK DRIVE, HOUSTON, TEXAS 77042

(713)975-0547   FAX (713)975-6338

Release: 2411-1D06

Revision C

Date: March 11, 1993

WARRANTY

Ruska Instrument Corporation warrants its products to conform to or exceed the specifications as set forth in its catalogs in use at the time of sale and reserves the right, at its own discretion, without notice and without making similar changes in articles previously manufactured, to make changes in materials, designs, finish, or specifications.  Ruska Instrument Corporation warrants products of its own factory against defects of material or workmanship for a period of one year from date of shipment.

Liability of Ruska Instrument Corporation under this warranty shall be limited to replacing, free of charge (FOB Houston, Texas), any such parts proving defective within the period of this warranty, but will not be responsible for transportation charges or consequential damages.

This warranty is not made for products manufactured by others which are illustrated and described in Ruska catalogs or incorporated in Ruska products in essentially the same form as supplied by the original manufacturer.  However, Ruska Instrument Corporation agrees to use its best efforts to have original suppliers make good their warranties.

COPYRIGHT NOTICE

Copyright ( 1989-1993 by Ruska Instrument Corporation. All rights reserved. This document may not be reproduced in part or in whole without the express written consent of Ruska Instrument Corporation.

DISCLAIMER

No representations or warranties are made with respect to the contents of this user's manual. Further, Ruska Instrument Corporation reserves the right to revise this manual and to make changes from time to time in the content hereof without obligation to notify any person of such revision.

TRADEMARK NOTICE

RUSKA ( is a trademark of Ruska Instrument Corporation.

Trademarks or tradenames are subject to state and federal laws concerning their unauthorized use or other infringements. The fact that the product marks or names in this manual do not bear a trademark symbol DOES NOT mean that the product name or mark is not registered as a trademark or tradename. Any queries concerning the ownership or existence of any trademarks or tradenames mentioned in this manual should be independently confirmed with the manufacturer or distributor of the product.

REVISION NOTICE

RELEASE

NUMBER

REV.

DATE OF

REVISION

DESCRIPTION

2411-1D00-00004

11/01/89

Original Release

2411-1D01-00004

05/14/90

Text and drawing changes .

ECO 15192M & 15068.

2411-1D06

A

09/24/90

Part No. Change - ECO 15459

2411-1D06

B

10/01/92

Text changes - ECO 18774

2411-1D06

C

03/11/93

Corrected equation Page 5-33 - ECO 18915

REVISION HISTORY

RELEASE 2411-1D00-00004 (11/01/89)

Original release

RELEASE 2411-1D01-00004 (05/04/90)

Section 2.3.2 Item 2.  Changed temperature range.

Section 2.3.3 Item 2g.  Renumbered from Item 3. Changed temperature range and added sentence.

Page C-3 No. 5.  Changed thickness of glass disk from .250" to .400".

Page C-5 No. 5.  Changed thickness of glass disk from .250" to .300".

Revised and updated drawings and Bills of Material.

RELEASE 2411-1D06 REVISION A (09/24/90)

Part number Change to conform to new numbering system - ECO 15459

RELEASE 2411-1D06 REVISION B (10/01/92)

Page 5-35 - changed 24 mo. Interval to 12 mo.

RELEASE 2411-1D06 REVISION C (03/11/93)

Corrected equation on page 5-33 - ECO 18915.

WARNING

PRESSURED VESSELS AND ASSOCIATED EQUIPMENT ARE POTENTIALLY DANGEROUS. THE APPARATUS DESCRIBED IN THIS MANUAL SHOULD BE OPERATED ONLY BY PERSONNEL TRAINED IN PROCEDURES THAT WILL ASSURE SAFETY TO THEMSELVES, TO SPECTATORS, AND TO THE EQUIPMENT.

TABLE OF CONTENTS

WARRANTY	i

COPYRIGHT NOTICE, DISCLAIMER, TRADEMARK NOTICE 	ii

REVISION NOTICE 	iii

REVISION HISTORY 	iv

WARNING 	v

TABLE OF CONTENTS 	vii

LIST OF FIGURES 	ix

SECTION 1: INTRODUCTION

1.1	SYSTEM OVERVIEW 	1-1

1.2	MANUAL OVERVIEW 	1-1

SECTION 2: AUTOPROMPT CONVERTER OVERVIEW

2.1	PURPOSE 	2-1

2.2	DESCRIPTION 	2-1

2.3	SPECIFICATIONS 	2-3

2.3.1	General 	2-3

2.3.2	Environmental 	2-3

2.3.3	Electrical 	2-3

2.3.4	Mechanical 	2-4

2.3.5	Calibration/Adjustments 	2-4

2.4	PRESSURE GENERATION AND MEASUREMENT  	2-5

2.4.1	Discussion of "Reference Plane" 	2-5

2.4.2	Pressure at Reference Plane of Device Under Test 	2-6

2.4.3	Absolute Pressure Reference 	2-7

2.4.4	Gage Pressure Reference 	2-7

2.4.5	Gas and Oil Heads 	2-8

2.4.6	Mass Equivalent of Pressure	2-8

2.4.7	Pressure Equivalent of Mass	2-9

2.4.8	Gas Medium Head Corrections for Gas Density	2-9

2.4.9	Oil Medium Head Corrections for Oil Density	2-11

2.5	MEASUREMENT OF TEMPERATURE STABILITY 	2-11

2.6	MEASUREMENT OF FLOAT POSITION	2-12

2.7	MEASUREMENT OF SINK RATE	2-12

2.8	'FLOATING' CRITERIA	2-12

2.9	DISCUSSION OF SYSTEM STABILITY	2-13

SECTION 3: USER INTERFACE

3.1	CAUTIONS 	3-1

3.2	DEFINITION OF TERMS 	3-1

3.3	DESCRIPTION OF THE KEY PANEL 	3-2

3.3.1	Keyboard and Key Functions 	3-2

3.3.2	Display 	3-2

3.4	OPERATION OF THE ENTRY FIELD EDITOR 	3-2

3.5	MENU ORGANIZATION 	  3-4

SECTION 4:  STARTUP AND CONFIGURATION

4.1	CONNECTING EXTERNAL COMPONENTS 	4-1

4.1.1	Connecting an RTD Sensor 	4-1

4.1.2	Connecting an FPI Sensor 	4-2

4.1.3	Connecting a Printer 	4-2

4.2	STARTUP 	4-2

4.2.1	Powering On 	4-2

4.2.2	Sign-On Display 	4-2

4.2.3	Self-test 	4-3

4.3	INITIAL CONFIGURATION OF PARAMETERS 	4-6

4.4	TURNING OFF THE AUTOPROMPTER 	4-7

SECTION 5: LOCAL OPERATION OF THE AUTOPROMPTER

5.1	ROUTINE OPERATION 	5-1

5.2	MAIN MENU OVERVIEW 	5-2

5.2.1	Main Menu #1 	5-2

5.2.2	Main Menu #2 	5-3

5.3	GENERATE PRESSURE (Accessed from Main Menu #1) 	5-3

5.3.1	Initialization Requirements 	5-3

5.3.2	Description of 'Generate Pressure' Function 	5-4

5.3.3	Entering the Desired Pressure 	5-5

5.3.4	Loading the Main and Gram Masses 	5-6

5.3.5	Achieving System Stability 	5-7

5.3.6	Acquiring a Reading from the Device Under Test 	5-8

5.3.7	Printing the Results 	5-9

5.4	MEASURE PRESSURE (Accessed from Main Menu #1) 	5-11

5.4.1	Initialization Requirements 	5-11

5.4.2	Description of 'Measure Pressure' Function 	5-11

5.4.3	Achieving System Stability 	5-12

5.4.4	Entering the Applied Mass 	5-13

5.4.5	Displaying the Measured Pressure 	5-14

5.4.6	Printing the Results 	5-14

5.5	SET PRESSURE UNITS (Accessed from Main Menu #1) 	5-15

5.5.1	Pressure Unit Screen #1 	5-16

5.5.2	Pressure Unit Screen #2 	5-16

5.5.3	Geopotential Feet of Altitude Equations 	5-17

5.5.4	Calibrated Air Speed (Knots) Equations 	5-17

5.6	DISPLAY PARAMETERS (Accessed from Main Menu #1) 	5-17

5.6.1	Display Piston/Cylinder Set's Data 	5-18

5.6.2	Display Mass Set's Data	5-19

5.6.3	Display Environment Parameters 	5-21

5.6.4	Display Date and Time 	5-23

5.7	SELECT PISTON/CYLINDER SET (From Main Menu #1) 	5-23

5.8	SELECT MASS SET (Accessed from Main Menu #1) 	5-24

5.9	SET PARAMETERS (Accessed from Main Menu #1 or #2) 	5-26

5.9.1	Configure Piston/Cylinder Set's Data 	5-27

5.9.2	Configure Mass Set's Data 	5-29

5.9.3	Set Environment Parameters 	5-31

5.9.4	Set Date and Time 	5-34

5.10	CALIBRATE SENSORS (Accessed from Main Menu #2) 	5-34

5.10.1	RTD Sensor Auto-cal Adjustment 	5-35

5.10.2	RTD Sensor Temperature Offset Calibration 	5-36

5.10.3	FPI Sensor Adjustment 	5-38

5.11	CONFIGURE SYSTEM (Accessed from Main Menu #2) 	5-42

5.11.1	Screen 1:  Installation Options	5-43

5.11.2	Screen 2:  Printer Configuration	5-44

5.11.3	Screen 3:  'A' Matrix Butterworth Filter Coefficients	5-45

5.11.4Screen 4:  'B' Matrix Butterworth Filter Coefficients	5-45

REFERENCES 		R-1

APPENDIX A:	AUTOPROMPT CONVERTER MENU FLOW DIAGRAM 	A-1

APPENDIX B:	EXAMPLES OF AUTOPROMPTER MAIN AND SUB MENUS	B-1

APPENDIX C:	AUTOPROMPTER INTERFACING TO DEADWEIGHT GAGES	C-1

APPENDIX D:	DRAWINGS AND BILLS OF MATERIAL 	D-1

APPENDIX E:	TEST AND CALIBRATION REPORTS 	E-1

LIST OF FIGURES

FIGURE 1-1:	PRESSURE GENERATION AND MEASUREMENT SYSTEM 	1-2

FIGURE 2-1:	MODEL 2411 AUTOPROMPT CONVERTER 	2-2

FIGURE 3-1:	AUTOPROMPT CONVERTER KEYBOARD 	3-2

FIGURE 4-1:	AUTOPROMPT CONVERTER BACK PANEL 	4-1

THIS PAGE INTENTIONALLY LEFT BLANK