# Purpose

RPTLoad version2 software is the replacement for RPTLoad version 1 (old DHI freeware) and is used to write or read Fluke “RPT” sensor characterization information stored within a product using the RS232 port.

# Data File

The characterization data files are ASCII delimited files containing the data used for the RPTLoad process. The file names are in the format SNserial.v01

# Communicating

Launch the program by double-clicking on the executable file “RPTLoad.exe”.  After the RPTLoad.exe utility is started, the next step is to configure the RS232 port settings by clicking on “Change” in the “Current Communication Settings” box.  Then select the computer’s RS232 port to use and the port settings.

The RS232 settings are 9600,N,8,1 for PM600 and PM630 modules.  Note that a PM600 or PM630 module must be installed in a calibration sled kit (not the pressure controller) to communicate with RPTLoad ver.2 software.

Once the RS232 settings have been configured then the communications can be tested from the Remote Communications tool which is located in the [Setup] drop-down menu, <Remote Communication>.    The Remote Communications tool allows commands to be sent and responses received from the PM600 module. In the command entry field, enter *0100SN and click on “Send”.  A valid response is the command sent echoed back with the serial number of the PM600’s transducer.  (Note – the serial number of the transducer is not the same as the serial number of the PM600 module assembly.)

## Retrieve Existing As-Found Characterization

Retrieving the existing characterization serves double purpose as a backup copy for the end-user and enables the Fluke Metrology team to validate the transducer information.  Click on the button “Download Coefficients” to begin the retrieval process.

A successful read of the coefficients opens a confirmation window, followed by a prompt to save the file.

The file will be saved by the sensor serial number and the “.v01” file extension.  It is important not to confuse this file with the new one which will be provided from the Fluke Metrology team because the new file will have the same file name!  Consider modifying the As-Found file name by appending it with the word “archived”.

## Upload New Characterization

A new “.v01” characterization file will be provided by the Fluke Metrology team. The preferred method of selecting the file to be uploaded is to use the [Setup] drop-down menu and select the option “Upload From File”.  Then navigate to the location where the “.v01” file is saved.  Be careful not to use the archived “As Found” coefficient file which was previously retrieved because the new file sent by Fluke uses the same file naming convention!  (The naming convention is a function the RPTLoad utility.)

When the “.v01” file is located, clicking “Open” will automatically start the write process to the PM600 module.

A successful write process message will appear at the bottom of the window when the process is complete.  Exit the RPTLoad program.

If a fault or an error occurs, then try the process a second time.  If the fault continues then the .v01 file may be corrupted.  Please contact Technical Support and request a new file.

If a file for the specified serial number does not exist, an error will occur.